# Annex No. 1 
 Technical Specification of the Supply

7 pieces of 6 channels data loggers + 10 pieces of 3 channels data loggers

| Mandatory parameters | Value | Parameter accomplished: Yes/No | Real value |
| --- | --- | --- | --- |
| Sample Rate | >= 5000 sps |  |  |
| Internal Data Storage | yes |  |  |
| Removable Data Storage | yes |  |  |
| Data Type | miniSEED |  |  |
| Streaming Data Format | SeedLink |  |  |
| Ethernet | yes |  |  |
| A/D converter | >= 24 bit |  |  |
| Environmental protection (IP) | >= 67 |  |  |
| Mass Centering/Locking/Unlocking | yes |  |  |
| Supply Voltage | 9 – 30 V |  |  |
| External GNSS antenna | yes |  |  |

| Rated parameters | Value | Parameter accomplished: Yes/No | Real value | Maximum Points |
| --- | --- | --- | --- | --- |
| Dynamic Range typical @ 100 sps | >= 140 dB |  |  | 10 |
| Size of internal Storage | >= 8 GB |  |  | 5 |
| Size of Removable Storage | >= 32 GB |  |  | 10 |
| SOH in miniSEED format | yes |  |  | 3 |
| Operating Temperature | <= -20 ˚C |  |  | 10 |
| Timing Accuracy | <= 10 us |  |  | 2 |
| Power Consumption | <= 1.5 W |  |  | 10 |
| VPN | yes |  |  | 5 |
| Orientation Correction | yes |  |  | 2 |

7 pieces of broadband seismometer

| Mandatory parameters | Value | Parameter accomplished: Yes/No | Real value |
| --- | --- | --- | --- |
| Symmetric Triaxial | yes |  |  |
| Force Balance System | yes |  |  |
| Sensitivity | <= 1000 V/m/s |  |  |
| Bandwidth | 10 sec to 100 Hz |  |  |
| Velocity Output | 40 V |  |  |
| Environmental protection (IP) | >= 67 |  |  |
| Supply Voltage | 9 - 30 V |  |  |

| Rated parameters | Value | Parameter accomplished: Yes/No | Real value | Maximum Points |
| --- | --- | --- | --- | --- |
| Clip Level | >= 20 mm/s |  |  | 10 |
| Maximum Operating Tilt | >= 5˚ |  |  | 5 |
| Dynamic Range typical @ 1Hz | >= 150 dB |  |  | 5 |
| Power Consumption | <= 0.5 W |  |  | 10 |
| Operating Temperature | <= -20 ˚C |  |  | 10 |
| Mass Centering required | no |  |  | 3 |

9 pieces of accelerometers

| Mandatory parameters | Value | Parameter accomplished: Yes/No | Real value |
| --- | --- | --- | --- |
| Symmetric Triaxial | yes |  |  |
| Force Balance System | yes |  |  |
| Maximum Full Scale Range | +-4 g |  |  |
| Minimum Full Scale Range | +-0.25 g |  |  |
| Bandwidth | DC to 200 Hz |  |  |
| Acceleration output | 40 V |  |  |
| Supply Voltage | 9 - 30 V |  |  |
| Environmental protection (IP) | >= 67 |  |  |

| Rated parameters | Value | Parameter accomplished: Yes/No | Real value | Maximum Points |
| --- | --- | --- | --- | --- |
| Dynamic Range typical @ 1Hz | >= 155 dB |  |  | 10 |
| Power Consumption | <= 1.5 W |  |  | 10 |
| Operating Temperature | <= -20˚C |  |  | 10 |

1 piece of accelerometer with integrated data logger

| Mandatory parameters | Value | Parameter accomplished: Yes/No | Real value |
| --- | --- | --- | --- |
| Symmetric Triaxial | yes |  |  |
| Force Balance System | yes |  |  |
| Maximum Full Scale Range | +-4 g |  |  |
| Minimum Full Scale Range | +-0.25 g |  |  |
| Bandwidth | DC to 200 Hz |  |  |
| Acceleration output | 40 V |  |  |
| Supply Voltage | 9 - 30 V |  |  |
| Environmental protection (IP) | >= 67 |  |  |
| Internal Data Storage | yes |  |  |
| Removable Data Storage | yes |  |  |
| Data Type | miniSEED |  |  |
| Streaming Data Format | SeedLink |  |  |
| Ethernet | yes |  |  |
| A/D converter | >= 24 bit |  |  |
| External GNSS antenna | yes |  |  |

| Rated parameters | Value | Parameter accomplished: Yes/No | Real value | Maximum Points |
| --- | --- | --- | --- | --- |
| Digitizer Dynamic Range typical @ 100 sps | >= 140 dB |  |  | 10 |
| Accelerometer Dynamic Range typical @ 1Hz | >= 155 dB |  |  | 10 |
| Size of internal Storage | >= 8 GB |  |  | 5 |
| Size of Removable Storage | >= 32 GB |  |  | 10 |
| SOH in miniSEED format | yes |  |  | 3 |
| Operating Temperature | <= -20 ˚C |  |  | 10 |
| Timming Accuracy | <= 10 us |  |  | 2 |
| Power Consumption | <= 3 W |  |  | 10 |
| VPN | yes |  |  | 5 |

Power consumption data logger + broadband seismometer + accelerometer

| Rated parameters | Value | Parameter accomplished: Yes/No | Real value | Maximum Points |
| --- | --- | --- | --- | --- |
| Power Consumption | <= 3 W |  |  | 5 |